VitalsPatch

A Wearable Multi-Patient Health Monitoring System

Draft Project Report

Maya Desai • Eeshani Shilamkar • Arnav Mohan

Bobby Taing • Nate Snyder •  Anushka Misra

Department of Electrical and Computer Engineering

EE 475 — Embedded Systems Capstone

University of Washington, Seattle, WA 98195

May 24, 2026

# Contents

# Abstract

VitalsPatch is a wearable health monitoring system designed for clinical and care environments where a single practitioner must keep watch over many patients at once. The device continuously measures heart rate, blood oxygen saturation (SpO2), body temperature, and motion, and detects falls and out-of-range vital signs in real time. Sensor data is collected on a wearable module, processed on an STM32F4 microcontroller, and forwarded to a Raspberry Pi 4 that hosts a centralized dashboard displaying live trends and prioritized alerts for multiple patients simultaneously.

This draft report documents the team's progress to date. It covers the motivation and market context, the three-module system architecture, the implementation of each module, and the results achieved through the Phase 1 and Phase 2 milestones. Later sections covering full-system validation, calibration, and final results are outlined here and will be completed as the remaining integration and testing work concludes ahead of the Phase 3 demonstration.

# 1. Team Roles and Responsibilities

The project is divided across six members so that the sensor, processing, communication, and interface subsystems can be developed in parallel. Each member owns a primary area while contributing to shared integration and testing.

| Member | Primary Responsibility |
| --- | --- |
| Eeshani Shilamkar | User interface design and overall project management |
| Arnav Mohan | Sensor integration and wireless communication |
| Maya Desai | Data processing and trend analysis on the STM32 |
| Nate Snyder | Sensor calibration and wearable assembly |
| Anushka Misra | Hardware integration and testing; sensor-side acquisition firmware |
| Bobby Taing | Alert logic development and system testing |

# 2. Introduction and Motivation

Continuous monitoring of patient vital signs is critical in hospitals, outpatient clinics, and elderly and nursing care facilities, where rapid response to a fall or a sudden change in vitals can directly affect patient outcomes. In these settings the value is not a daily wellness summary but uninterrupted observation of many patients at the same time, with immediate notification when something goes wrong.

The wearable medical devices market reflects strong demand for this kind of technology. It was valued at approximately USD 42.74 billion in 2024 and is projected to reach USD 168.29 billion by 2030, a compound annual growth rate of about 25.53 percent. Despite this growth, most existing devices focus on only one or two measurements, which forces care facilities to combine several products to obtain a complete picture of a patient.

## 2.1 Market Gap and Comparable Products

Several wearable products address parts of this problem but none combines comprehensive multi-signal sensing with centralized, multi-patient monitoring:

- FloPatch (Flosonics Medical) — a clinical-grade wearable ultrasound device for assessing blood flow.

- Aktiia — a cuffless blood pressure monitoring bracelet.

- EvieMed Ring (Movano Health) — a ring that tracks pulse oximetry.

VitalsPatch differentiates itself by integrating heart rate, SpO2, temperature, and accelerometer- and gyroscope-based fall detection into a single wearable, and by pairing that wearable with a centralized dashboard that lets one healthcare professional monitor several patients from a single screen.

# 3. System Design and Architecture

The system is organized into three primary modules that together form an end-to-end pipeline from the patient to the monitoring dashboard:

Sensor / Data Acquisition  →  (Wireless)  →  STM32 Processing  →  (Wired Serial)  →  Raspberry Pi UI

- The sensor module collects physiological and motion data from the wearable and formats it into structured packets.

- Data is transmitted wirelessly to the STM32F4, which parses the packets and analyzes them for abnormal conditions.

- The STM32 forwards processed data and alert codes to the Raspberry Pi 4 over a wired serial connection.

- The Raspberry Pi hosts a dashboard that displays near real-time trends and prioritized alerts for each monitored patient.

## 3.1 Hardware Components

The table below summarizes the role of each major component in the system.

| Component | Role | Interface |
| --- | --- | --- |
| MAX30102 | Heart rate and SpO2 sensing (optical PPG) | I2C → sensor MCU |
| MPU-6500 | Motion and fall detection (accelerometer + gyroscope) | I2C / SPI → sensor MCU |
| Temperature sensor | Skin / body temperature sensing | ADC / I2C |
| ESP32-C3 | Sensor-side acquisition and packet transmission | Wireless → STM32F4 |
| STM32F4 | Data processing, event detection, alert codes | Serial → Raspberry Pi |
| Raspberry Pi 4 | Receiver, dashboard host, logging | Local dashboard software |
| LiPo battery | Portable power for the wearable | Regulated 3.3 V rail |

# 4. Implementation

Each of the three modules was developed to function independently before the team integrated them into a single pipeline.

## 4.1 Sensor and Data Acquisition Module

The wearable module reads heart rate and SpO2 from the MAX30102 and motion data from the MPU-6500. The sensor-side microcontroller polls both sensors at a consistent sampling rate, populates a structured data record, and transmits it as framed packets to the STM32. Framing the data with a fixed header and checksum keeps transmission reliable and lets the receiver detect dropped or corrupted packets.

- Reads physiological and motion data at a consistent sampling rate.

- Formats readings into structured transmission packets.

- Transmits packets to the STM32 for processing.

## 4.2 Data Processing Module (STM32F4)

The STM32F4 receives sensor packets, validates them, and runs threshold-based analysis to detect abnormal conditions. When a reading crosses a clinically informed threshold the STM32 raises an alert code, where 0 indicates a normal state. The processed output is a set of recent data arrays paired with the corresponding alert codes.

| Code | Detected Event |
| --- | --- |
| 0 | Normal — all readings within range |
| 1 / 2 | Heart rate above / below threshold |
| 3 / 4 | Temperature spike / drop |
| 5 | Low SpO2 |
| 6 | Fall detected via accelerometer / gyroscope |

Example processed packet:  HR:[140,138,142]; SPO2:[85,86,84]; TEMP:[102.5,103.0,102.8]; ERR:[1,2,3,4,5,6]

## 4.3 User Interface Module (Raspberry Pi 4)

The Raspberry Pi hosts a dashboard that visualizes incoming data and surfaces alerts. Time-series graphs update in near real time for heart rate, temperature, and SpO2, while alerts are given visual priority over trend graphs so that a critical event is immediately obvious. The interface keeps a per-patient log of recent alerts and events and updates dynamically as new data arrives.

- Near real-time time-series graphs for heart rate, temperature, and SpO2.

- Alerts displayed with higher visual priority than graphs.

- Separate, clearly distinguished views and logs for each patient.

# 5. Development Progress and Results

## 5.1 Phase 1 — Independent Modules

Phase 1 focused on building and validating each module on its own and establishing initial communication between the sensor module and the STM32. By the end of Phase 1 the team had achieved:

- A sensor module that reads all required signals and formats them into transmission packets.

- An STM32 capable of processing data and generating the correct alert codes for each event type.

- A Raspberry Pi dashboard displaying near real-time graphs and prioritized alerts on mock data.

- Functional communication from the sensor module to the STM32.

The dashboard was verified against a series of test packets covering a baseline state, each individual alert in isolation, and combined emergency scenarios in which multiple alerts fire at once.

## 5.2 Phase 2 — Full Pipeline and Multi-Patient Support

Phase 2 connects the modules into a complete pipeline and extends the dashboard to monitor two patients at the same time. Current status:

- Working: near real-time graphs for two patients, clear separation of Patient 1 and Patient 2 data, prioritized alerts, per-patient logs, and dynamic updates as new data arrives.

- Working: two independent sensor modules each collecting and transmitting data with patient identification.

- Finalizing: the wired serial link delivering processed data from the STM32 into the Raspberry Pi UI, targeted for completion by the end of the week.

# 6. Testing and Validation

# 7. Risks and Contingency

Key risks identified during planning, with planned mitigations, are summarized below; this section will be expanded with results once validation is complete.

- Pulse oximetry accuracy can vary with skin tone and motion — mitigated through calibration and by discarding readings flagged as motion-corrupted by the accelerometer.

- Routine activity such as coughing or sitting up may trigger false alerts — mitigated by requiring an abnormal reading to persist briefly before an alert fires.

- A wireless dropout could take a patient offline silently — mitigated by periodic heartbeat signals so the dashboard knows each device is still online.

- A dashboard crash could lose recent data — mitigated by storing recent readings locally on the wearable as a backup.

# 8. Future Work

# 9. Conclusion

# 10. References

[1] Grand View Research, “Wearable Medical Devices Market Size & Trends.” Available: grandviewresearch.com

[2] National Science Foundation, America’s Seed Fund (SBIR/STTR), Medical Devices topic. Available: seedfund.nsf.gov

[3] Seeed Studio, XIAO ESP32-C3 Getting Started Guide. Available: wiki.seeedstudio.com